# Extended Hydrogen Sensor Performance Comparison (Draft)

DRAFT -- not part of the approved knowledge base. This extends the approved “Competitor Comparison_Hydrogen Sensor Performance” document with two additional real competitors (H2scan and Figaro Engineering), researched from their own public datasheets/product pages. It also grounds MNST's figures in a specific, currently approved product -- the AURIGA Portable H2 Leak Detector catalog -- rather than the original document's more abstract “MEMS/SSEC platform” framing, per instruction to derive the MNST side from the approved catalogs. Every figure below is either quoted from a cited source or explicitly marked “Not disclosed” where the source didn't cover it -- nothing here is invented or estimated.

## 1. Response Time

| Company / Sensor | Speed |
| --- | --- |
| Infineon XENSIV TCI | ≤100 ms |
| MNST AURIGA (Portable H2 Leak Detector) | T10 <0.5 s, T90 <3 s |
| Honeywell HLD | <2 seconds |
| Dräger XXS H2 HC | ≤20 seconds |
| H2scan HY-ALERTA 5020/5021 | 5020: <60 s to report 2% (>3% applied); 5021: T90 5 minutes |
| Figaro TGS2616-C01 | Not disclosed on the datasheet reviewed (only sensitivity-vs-concentration curves given) |
| Honeywell Sensepoint XCD | Not disclosed |

## 2. Detection Range

| Company / Sensor | Range |
| --- | --- |
| Honeywell Sensepoint XCD | Up to 1,000 ppm |
| Figaro TGS2616-C01 | 30–3,000 ppm (typical detection range, per datasheet) |
| Infineon TCI | 0–160,000 ppm |
| MNST AURIGA | 0–5,000 ppm / 0–5 vol% / 0–10 vol% (selectable ranges); Minimum Detection Limit 10 ppm |
| Dräger XXS H2 HC | 0–40,000 ppm |
| Honeywell HLD | 0–40,000 ppm |
| H2scan HY-ALERTA 5020/5021 | 0.4–5% H2 (10–125% LEL) = 4,000–50,000 ppm |

## 3. Durability / Service Life

| Company / Sensor | Lifetime |
| --- | --- |
| Infineon TCI | 15-year rating |
| Honeywell HLD | “Minimal maintenance over a decade” |
| MNST AURIGA | Operating life of 5 years (per product features); single-shot calibration recommended every 6 months, mandatory every 12 months |
| Dräger XXS H2 HC | More than 2 years |
| H2scan HY-ALERTA | Self-calibrating technology, “None” required per datasheet -- no stated end-of-life figure found |
| Figaro TGS2616-C01 | Not disclosed on the datasheet reviewed |
| Honeywell Sensepoint XCD | Not disclosed |

## Sensing Principle

| Product | Principle |
| --- | --- |
| Infineon TCI | Thermal conductivity |
| Honeywell HLD | Thermal conductivity |
| MNST AURIGA | Solid-state electrochemical (solid electrolyte, patented reference electrode design) |
| Dräger XXS H2 HC | Electrochemical |
| Honeywell Sensepoint XCD | Electrochemical |
| H2scan HY-ALERTA | Proprietary hydrogen-specific resistive sensor (solid-state, hydrogen-selective by design) |
| Figaro TGS2616-C01 | Metal-oxide semiconductor (MOS) |

## 4. Operating Environment & Precision

| Company / Sensor | Operating Temp. | Humidity | Pressure | Warm-up / Response | Accuracy / Precision |
| --- | --- | --- | --- | --- | --- |
| Honeywell HLD | -40°C to 85°C | 10–90% RH | 80–110 kPa | <1 second | ±10% (stated for -40°C to 65°C only) |
| Honeywell Sensepoint XCD | Not disclosed | Not disclosed | Not disclosed | Not disclosed | Not disclosed |
| Infineon TCI | -40°C to 105°C | 0–100% RH (no condensation) | 50–130 kPa | ~20 ms (typ.) | ±0.12 vol%H2 with pressure comp.; ±0.2 vol%H2 without |
| Dräger XXS H2 HC | -20°C to 50°C | 10–90% RH | 700–1,300 hPa | ≤1 hour | ≤±2% of measured value; zero drift ≤±0.05 vol%/year |
| MNST AURIGA | -20°C to +65°C | Up to 99% RH at 40°C, non-condensing (sensor probe) | 1 Bar to 1.2 Bar | T10 <0.5 s, T90 <3 s | 0–10ppm ±5ppm; 10–100ppm ±20ppm; 100–1000ppm ±100ppm; 1000–5000ppm ±5% MV; 0.5–10% ±2% MV |
| H2scan HY-ALERTA 5021 | -40°C to +60°C | <95% RH (non-condensing) | Not disclosed | T90 5 minutes | ±0.3% H2 (accuracy and repeatability) |
| Figaro TGS2616-C01 | Not disclosed as an operating range (test condition only: 20±2°C, 65±5%RH) | See temp. column | Not disclosed | Not disclosed | Sensitivity expressed as resistance ratio (Rs), not a ±% accuracy figure -- not directly comparable to the other rows |

## 5. Cross-Sensitivity / Interference Notes

“Cross-sensitivity” means a sensor reacts to a gas other than the one it's meant to measure, producing a false or skewed reading. This matters in real environments where more than one gas may be present at once.

| Company / Sensor | Gas / Interferent | Effect on Reading |
| --- | --- | --- |
| Honeywell HLD | Helium; Carbon dioxide | Affected above 1% CO2 (0.04% CO2 is considered the normal range); helium also listed as cross-sensitive |
| Dräger XXS H2 HC | Carbon monoxide (1,000 ppm) | False reading of ≤0.1 vol% H2 |
| Dräger XXS H2 HC | Nitrogen monoxide (20 ppm) | False reading of ≤0.05 vol% H2 |
| Dräger XXS H2 HC | Hydrogen sulfide; sulfur dioxide | Eliminated by an internal selective filter (no effect on reading) |
| MNST AURIGA | Other combustible and reducing gases | No cross sensitivity per datasheet; selective and specific to H2 only |
| Figaro TGS2616-C01 | Acetylene (C2H2) | “Small interference from acetylene” -- newly developed sensing element specifically reduces this; selectivity ratio given as Rs(500ppm C2H2)/Rs(3ppm H2) > 1 |
| Infineon TCI | Not disclosed | Not disclosed |
| Honeywell Sensepoint XCD | Not disclosed | Not disclosed |
| H2scan HY-ALERTA | Not disclosed on the page reviewed | Not disclosed |

## 6. Certifications & Hazardous-Area Approval

Certifications differ in scope: some cover use in explosive (hazardous) atmospheres, others cover automotive qualification, general product safety, or nothing beyond a component-level mark.

| Company / Sensor | Certifications Cited | Scope |
| --- | --- | --- |
| Honeywell HLD | RoHS; REACH; IP67; ASIL QM | Automotive/environmental standards only (ISO 11452, ISO 16750, IEC 60068 series); no ATEX/IECEx hazardous-area certification disclosed |
| Honeywell Sensepoint XCD | ATEX; IECEx; UL; CSA; INMETRO; KTL; GB; CCCF; IP66 | Certified for use in explosive (hazardous) atmospheres across multiple global standards |
| Infineon TCI | AEC-Q100 Grade 2 (automotive); JEDEC JESD47L, JESD22, J-STD-020 (industrial) | Component-level qualification; not a hazardous-area detector certification |
| Dräger XXS H2 HC | Not disclosed | Specific ATEX/IECEx certificate number not found on the pages reviewed |
| MNST AURIGA | None disclosed in the catalog; IP67 ingress protection stated | Catalog explicitly states “Area of Deployment: Safe Area. This instrument should not be deployed in a Hazardous Area.” -- i.e. this specific product is documented as NOT intended for hazardous-area use |
| H2scan HY-ALERTA 5020/5021 | C/US, CE 0359 (per product page) | General product safety/EMC marks; the page reviewed did not state an ATEX/IECEx hazardous-area rating |
| Figaro TGS2616-C01 | Not disclosed on the datasheet reviewed | Component-level sensor; no certification section on this datasheet |

## 7. Power Consumption / Electrical Supply

| Company / Sensor | Supply | Current / Power Draw |
| --- | --- | --- |
| Honeywell HLD | 8 to 36 Vdc | 16 mA @ 12 Vdc typ.; 8 mA @ 24 Vdc typ.; in-rush 100 mA |
| Honeywell Sensepoint XCD | Not disclosed | Not disclosed |
| Infineon TCI | 3.3 V ± 5% | Stand-by 2.3–6 µA; peak 5–7 mA; typ. 112 µA at 1 measurement/second |
| Dräger XXS H2 HC | Not disclosed | Not disclosed |
| MNST AURIGA | 3x rechargeable AA Ni-MH batteries; USB C-Type charging | Battery life 12–24 hours per full charge, depending on alarm/detection frequency |
| H2scan HY-ALERTA 5020/5021 | 9–48 VDC | 10 W |
| Figaro TGS2616-C01 | Heater 5.0±0.2 V DC; Circuit 5.0±0.2 V DC | Heater current 56±5 mA; heater power 280 mW |

## 8. Physical Size & Weight

Form factor matters differently depending on whether the product is a finished, housed detector, a handheld instrument, or a bare component intended for integration.

| Company / Sensor | Dimensions | Weight |
| --- | --- | --- |
| Honeywell HLD | 72.63 × 77.5 × 17.38 mm | 50 g |
| Honeywell Sensepoint XCD | Not disclosed | Not disclosed |
| Infineon TCI | 14-pin PG-DSOSP-14-84 package (bare component) | Not disclosed (die-level component) |
| Dräger XXS H2 HC | Miniature “XXS” cartridge; exact dimensions not disclosed on pages reviewed | Not disclosed |
| MNST AURIGA | Monitor: 145 × 88 × 34 mm; Probe: 260 mm length, ø24mm handle OD, up to 1500mm connecting cable | Monitor ~300 g with batteries; up to ~400 g total with probe |
| H2scan HY-ALERTA 5020/5021 | 161 × 40 × 40 mm | Not disclosed |
| Figaro TGS2616-C01 | TO-5 metal can, ø9.2mm, ~10mm height (bare component) | Not disclosed |

## 9. Communication / Output Interface

This is how each sensor reports its reading to a controller, display, host system, or user.

| Company / Sensor | Interface |
| --- | --- |
| Honeywell HLD | CAN (SAE J1939, 250 kBaud); analog voltage and PWM output variants available on request |
| Honeywell Sensepoint XCD | 4–20 mA (sink/source selectable); MODBUS |
| Infineon TCI | I2C digital interface, up to 1 Mbit/s |
| Dräger XXS H2 HC | Not disclosed (analog electrochemical output via host instrument electronics) |
| MNST AURIGA | USB C-Type for computer interface (data transfer, graphs, diagnostics, device management); OLED on-device display; no RS485/4-20mA/CAN documented for this specific handheld product |
| H2scan HY-ALERTA 5020/5021 | RS485 (Modbus RTU, 19200 baud); optional 4-20 mA via HYAO-1 accessory; dry-contact relay outputs |
| Figaro TGS2616-C01 | Analog resistive output only (raw sensor resistance via a simple external circuit) -- a bare sensing element, not a digital interface |

## References

MNST (all AURIGA figures): MNST_NC11. Catalogue_Auriga (Leak Detector Series).pdf -- approved document, data/approved_docs/

Infineon TCI: https://www.infineon.com/assets/row/public/documents/24/49/infineon-tci-datasheet-en.pdf

Honeywell HLD: https://prod-edam.honeywell.com/content/dam/honeywelledam/sps/siot/ja/products/sensors/electrification-sensors/hydrogen-safety-sensors/hld-series/documents/hon-ia-hss-hldhydrogen-leak-detection-datasheet-en.pdf

Honeywell Sensepoint XCD: https://prod-edam.honeywell.com/content/dam/honeywell-edam/sps/his/en-us/products/gasand-flame-detection/documents/ds01082_sensepoint_xcd_brochure_flr_1_30_14.pdf

DrägerSensor XXS H2 HC: https://www.draeger.com/Content/Documents/Content/XXS-H2-HC-6812025.pdf

H2scan HY-ALERTA 5020/5021: https://h2scan.com/product/hyalerta-hydrogen-monitor/

Figaro TGS2616-C01: https://www.figarosensor.com/product/docs/tgs2616-c01_product%20information(fusa)rev00.pdf